## Sheet1
DO NOT MODIFY THE WORK SHEET. FILL UP THE BLUE CELLS ONLY
Name of Medical Physicist responsible for the measurements
Date of completion of the measurements
1 | General information of the Simulator Units
1.1 | Type of the unit (CT-SIM)
1.2 | Manufacturer
1.3 | Model
1.4 | Sr. No.
1.5 | Bore diameter in cm
1.6 | Table top surface : (flat/curved)
1.7 | Maximum Tube Voltage in kVp
1.8 | Maximum Tube Current in mA
1.9 | Available slice thicknesses in mm
1.1 | Scan time per rotation
1.11 | Max. No. of slices per rotation
1.12 | Scan time for full helical scan
1.13 | Focal spot size
1.13.1 | Small |  | ___ mm x ___ mm
1.13.2 | Large |  | ___ mm x ___ mm
1.14 | Type of Rectification (Full Wave/ Half Wave/ Multi pulse/ Other (specify))
1.15 | Filtration
1.15.1 | Inherent (mm of Al)
1.15.2 | Added (mm of Al)
1.15.3 | Total (mm of Al)
1.16 | Type of rotation mechanism (slip ring/ other)
1.17 | Gantry tilt (maximum)
1.18 | Focus-isocenter distance
1.19 | Maximum Field of View (FOV)
1.2 | Maximum movement of the scanning table
1.21 | Minimum table increment available for scan
1.22 | Type of detector (solid state/ gas filled/any other specify)
1.23 | Number of detectors
2 | Performance test results of the CT Simulator unit
Sr. No. | Parameters | Tolerance Value, if any | Observation (Write Yes/No or give values of measurment as the case may be) | Observation is meeting the requirments | Remarks, if any
2.1 | Controls on the Gantry provided and working satisfactorily
2.2 | Emergency switch provided and working satisfactorily
2.3 | Congruence between the gantry midline and table mid line | ±5 mm
2.4 | Gantry tilt accuracy | ±1°
2.5 | Accuracy of longitudinal patient support positioning | ±1mm
2.6 | Linearity of Table vertical and | ±1 mm over the
| longitudinal motion | range of table
motion
2.7 | Reproducibility of Table vertical and longitudinal motion | ±1 mm over the
range of table
motion
2.8 | Alignment of gantry lasers with the center of imaging plane | ±2 mm
2.9 | Orientation (parallelism and orthogonality) of gantry lasers with respect to the imaging plane | ±2 mm over the
length of laser
projection
2.1 | Spacing of lateral wall lasers with | ±2 mm
respect to lateral gantry lasers and scan plane
2.11 | Orientation (parallelism and orthogonality) of wall lasers with respect to the imaging plane | ±2 mm over the
length of laser
projection
2.12 | Linearity of moving Laser | ±1 mm
2.13 | Reproducibility of moving Laser | ±1 mm
2.14 | Table top sag at isocenter when loaded with 135kg distributed over 2m through isocentre | 2 mm
2.15 | Slice thickness accuracy (sensitivity profile width) | ±1mm for slice thickness > 2mm
2.16 | Accuracy of kVp | ≤10%
2.17 | Accuracy of mA | ≤20%
2.18 | Accuracy of timer/irradiation time | ≤10%
2.19 | Accuracy of current time product (mAs) | ≤10%
2.2 | Linearity of timer (Coeff. of Linearity) | Coeff. of Linearity ≤ 0.1
2.21 | Linearity of tube current (Coeff. of Linearity) | Coeff. of Linearity ≤ 0.1
2.22 | Linearity of Output (Coeff. of Linearity) | Coeff. of Linearity ≤ 0.1
2.23 | Output reproducibility | Coeff. of Variation <0.05
2.24 | Focal spot size (f)
2.24.1 | for f< 0.8mm | 0.5f
2.24.2 | for 0.8 mm £ f £ 1.5 mm | 0.4f
2.24.3 | for f > 1.5 mm | 0.3f
2.25 | Total filtration (mm of Al equivalent) of the x-ray tube
2.25.1 | for kV ≤ 70 kV, | 1.5 mm Al
2.25.2 | for 70 < kV≤ 100 | 2 mm Al
2.25.3 | for kV> 100 | 2.5 mm Al
2.26 | Radiation leakage level from the tube at 1m | ≤ 1 mGy in 1 hr
2.27 | CT dose index (CTDIW ) | ±20% of specified value
2.28 | CT number uniformity | ±4 HU for center and ±2 HU for other location
2.29 | CT number Accuracy | ±4 HU
2.3 | Low contrast resolution | 5.0mm at 1% contrast difference (minimum)
2.5mm at 0.5% contrast diffence (expected)
2.31 | Spatial resolution | The size of the bar/hole pattern 0.5lp/cm should be resolved for 10%contrast diff. OR
10% of specified nominal value whichever is greater
2.32 | Noise | ±15% of the specified value
2.33 | Geometric Accuracy & Distortion (Verify the dimensions and spatial fidelity of known geometrical phantoms in the axial CT images) | ±2 mm
2.34 | Image Orientation (e.g. supine/prone, Superior-inferior, anterior- posterior and Left-right) | No difference
2.35 | Check scale, orientation, labeling and co-ordinates in sagittal, coronal and oblique images after image reconstruction | Same as axial slices
2.36 | Display of 2-D images and 3D visualization of multiple images | Visual check
2.37 | Verify accuracy of grayscale values, particularly for conversion of CT number to electron density | ±3% of user calculated electron density of the material
3 | Safety aspects of CT Simulator installation
3.1 | Radiation Warning legend and symbol at the entrance provided: (Yes/No)
3.2 | Radiation warning light at the entrance provided and working: (Yes/No)
3.4 | Whether entrance doors lined with 2mm lead are provided: (yes/No)
3.5 | Whether viewing window with 2 mm lead equivalence is provided: (yes/No)