## Sheet1
DO NOT MODIFY THE WORK SHEET. FILL UP THE BLUE CELLS ONLY
Name of Medical Physicist responsible for the measurements
Date of completion of the measurements
1 | General information of the kV Imaging system
1.1 | Model of accelerator on which it is mounted
1.2 | Manufacturer
1.3 | Model
1.4 | Sr. No.
1.5 | Imaging capabilities (Radiography, Fluroscopy, CBCT)
1.6 | Imaging device (Image intensifier tube (IIT) / flat panel)
1.7 | Focus to axis distance (FAD) in cm
1.8 | Maximum Tube Voltage in kVp
1.9 | Maximum Tube Current in mA
1.1 | Type of Anode (Rotating / Stationary)
1.11 | Focal spot size (Small: ___ mm x ___ mm & Large: __ mm x __mm)
1.12 | Type of Rectification (Full Wave / Half Wave/ Multi pulse/Other (specify))
1.13 | Filtration in mm of Al (Inherent: ________, Added:______ Total:_____)
2 | Performance test results of the kV Imaging system (*Applicable for cone beam CT)
Sr. No. | Parameters | Tolerance Value, if any | Observation (Write Yes/No or give values of measurment as the case may be) | Observation is meeting the requirments | Remarks, if any
2.1 | Collision interlocks provided and working
2.2 | Accuracy of Focal spot size (f)
2.2.1 | for f< 0.8mm | 0.5f
2.2.2 | for 0.8 mm £ f £ 1.5 mm | 0.4f
2.2.3 | for f > 1.5 mm | 0.3f
2.3 | Accuracy of kVp | ±5kVp
2.4 | Accuracy of timer | 10%
2.5 | Linearity of mA station | CoL≤0.1
2.6 | Linearity of timer | CoL ≤0.1
2.7 | Output consistency | CoV ≤0.05
2.8 | Total filtration (mm of Al equivalent) of the x-ray tube
2.8.1 | for kV ≤ 70 kV, | 1.5 mm Al
2.8.2 | for 70 < kV≤ 100 | 2 mm Al
2.8.3 | for kV> 100 | 2.5 mm Al
2.9 | ON Position Radiation Leakage level from X-Ray tube housing measured at maximum tube rating | < 1mGy in one hour at 1 m
2.1 | Positional/re-positional accuracy of imager | ≤ 1 mm
2.11 | Coincidence of centre of imaging system and isocentre (vertical gantry) | ≤ 1 mm
2.12 | Geometric distortion* | ≤ 1 mm
2.13 | Low Contrast resolution | Manufacturer’s specified tolerance |  |  | Manufacturer’s specified value______
2.14 | Spatial (high contrast) resolution | Manufacturer’s specified tolerance |  |  | Manufacturer’s specified value______
2.15 | Uniformity and noise | Manufacturer’s specified tolerance |  |  | Manufacturer’s specified value______
2.16 | Accuracy of HU values* | Manufacturer’s specified tolerance |  |  | Manufacturer’s specified value______
2.17 | Constancy of HU values* | Manufacturer’s specified tolerance |  |  | Manufacturer’s specified value______