## HDR QA Sheet
DO NOT MODIFY THE WORKSHEET. FILL UP THE BLUE CELLS ONLY.
Name of the Medical Physicist responsible for the measurements:
Date of completion of the measurements:
1 | General Information of the Brachytherapy unit:
1.1 | Type of the unit (LDR /MDR /HDR /PDR)
1.2 | Manufacturer
1.3 | Model
1.4 | Sr. No.
1.5 | Radionuclide (Cs-137 /Ir-192 /Co-60/ any other)
1.6 | No. of sources
1.7 | Dummy source provided (Yes/No)
1.8 | Maximum capacity of the unit (GBq) / (Ci )
1.9 | Source Strength (As per source certificate)
1.9.1 | Apparent Activity (GBq) / (Ci ) as on date
1.9.2 | Reference Air Kerma Rate (µGy/hr at 1metre distance) as on date
1.10 | Measured Source Strength
1.10.1 | Apparent Activity (GBq) / (Ci ) as on date
1.10.2 | Reference Air Kerma Rate (µGy/hr at 1metre distance) as on date
1.11 | Maximum no. of channels available
1.12 | Maximum No. of channels used simultaneously
1.13 | Source dimensions
1.13.1 | Physical Dimension [Length(mm), Dia (mm)]
1.13.2 | Active Dimension [Length(mm), Dia (mm)]
1.14 | Method of source transfer between applicator & source storage (Pneumatic / Motor driven/Any other)
1.15 | Treatment mode (Programmable/ Pulsed /Any other)
1.16 | Method of source retraction in the event of power failure (Motor with Back-up Battery/ Pneumatic /Any Other)
2 | Performance test results of the Remote Afterloading Brachytherapy unit
Sr. No. | Parameters | Tolerance Value, if any | Observation (Write Yes/No or give values of measurment as the case may be) | Observation is meeting the requirments | Remarks, if any
2.1 | Treatment Unit
2.1.1 | Coupling between guide tube and unit satisfactory
2.1.2 | Mechanism for locking the guide tube with unit (such as indexer locking ring) is provided and working
2.1.3 | Interlock for coupling of guide tube and Treatment unit are provided and working
2.1.4 | Coupling between guide tube and applicator satisfactory
2.1.5 | Interlock for coupling of guide tube and Applicator are provided and working
2.1.6 | Emergency stop button to interrupt the irradiation is provided and working
2.1.7 | Key operated locking mechanism is provided and working
2.1.8 | Source "OFF" and "ON" indicators are provided and working
2.1.9 | Whether guide tube is free of kinks
2.1.10 | Time taken to drive the source to “ON” position | manufacturer stated value:____ sec
2.1.11 | Time taken to drive the source to “OFF” position | manufacturer stated value:____ sec
2.1.12 | Interlock for treatment room door are provided and working
2.1.13 | Whether warning lights (‘OFF”, “ON” and “Transit” indicators) at/or near the entrance to the treatment room are provided and working
2.2 | Treatment control console
2.2.1 | Power switch is provided and working
2.2.2 | Door interlock indicator provided and working
2.2.3 | Treatment and elapsed/ remaining time display counters provided and working
2.2.4 | Indicator lights are provided and working:
2.2.4.1 | Ready state/Beam OFF – Green
2.2.4.2 | Source in transit- Yellow (flashing)
2.2.4.3 | Radiation Beam ON- Yellow
2.2.4.4 | Immediate action required to terminate an unintended status of operation- Red
2.2.4.5 | EQUIPMENT switched on but all the operations required to reach the READY STATE not yet carried out- White
2.2.5 | Treatment parameter reset switch is provided and working
2.2.6 | Treatment “ON” switch is provided and working
2.2.7 | Treatment “OFF” switch is provided and working
2.2.8 | Treatment “interruption(pause)” switch is provided and working
2.2.9 | Emergency switch is provided and working
2.2.10 | Whether display for set, elapsed / remaining time are retained in case of interruption/ termination until reset to zero
2.2.11 | Whether record of number of source transit cycle is displayed or retrievable
2.2.12 | Whether record of number of inactive (dummy) source cycle is displayed or retrievable
2.2.13 | Whether initiation & continuation of irradiation are possible only by an operation at control panel
2.2.14 | Whether removable key (or tool) for enabling operation at control console is provided and working
2.2.15 | Whether equipment is rendered inoperative on removal of the tool or key
2.2.16 | Whether additional key or designated password for confirmation of pre-set treatment parameter is provided and working
2.2.17 | Whether display of selection of channels, radioactive sources positions and movements of radioactive sources in each channel are provided
2.2.18 | Whether pre-set of dwell time in the specified range is possible
2.2.19 | Whether timer is able to control the pre-determined dwell position/time
2.2.20 | Whether pre-set or confirmation of dwell position/time are possible only at control console
2.2.21 | Whether timer stops temporarily during interruption
2.2.22 | Whether irradiation stops automatically after completion of treatment time
2.2.23 | Whether initiation is possible only when at least one dwell time has been pre-set to a value greater than zero
2.2.24 | Whether no change is possible in the pre-set value after initiation of irradiation without inducing or reaching termination
2.2.25 | Whether in case of termination, records of irradiation such as patient ID, source strength, date and time of initiation, termination, interruption, continuation, etc is available
2.3 | Emergency and failures
2.3.1 | In the event of unplanned power failure, whether battery-powered back-up power supply is provided to withdraw the source
2.3.2 | In emergency, whether mechanism for manual retraction of source is provided
2.3.3 | In case of failure of Timer, whether retrieval of total duration of irradiation at any dwell position is possible
2.3.4 | Whether source return automatically to the unit and indicators (light or other means) are activated in case of failure of
2.3.4.1 | - power supply
2.3.4.2 | - controlling timer
2.3.4.3 | - sources to be in their intended positions
2.3.4.4 | - source drive mechanism
2.3.4.5 | - interlocks
2.3.5 | Whether equipment is provided with a radiation monitor to detect failure of source to return to the unit
2.3.6 | Whether retrieval of preset parameters for a period of at least 10 hrs is possible, in case of any failure
2.3.7 | Whether visual indicators/message activated in case of failure of controlling timer, positional accuracy and movement of sources
2.3.8 | Whether no initiation or continuation of treatment is possible in the event of failure of controlling timer, positional accuracy and movement of sources
2.3.9 | Whether audible alarm and visible indication initiated immediately at the control console, in case of failure of source to arrive at its intended position
2.3.10 | Whether visual indicators are activated when initiation is prevented or irradiation is terminated due to any fault conditions
2.4 | Source Position Accuracy and Reproducibility
2.4.1 | Source position accuracy | ± 1 mm
2.4.2 | Source position reproducibility | ± 1 mm
2.5 | Performance of Timer
2.5.1 | Timer accuracy (for set time ≥ 1% of maximum pre-settable value) | ± 1 %
2.5.2 | Timer Linearity | ± 1 %
2.5.3 | Timer error | ± 1 %
2.6 | Source Strength Verification
2.6.1 | Deviation between specified and measured source strength | ± 3%
2.7 | Leakage and contamination test
2.7.1 | Contamination from the source | 200 Bq
2.7.2 | Maximum leakage radiation level at 5 cm from the surface of the HDR unit in source 'OFF’ condition | ≤ 100 µSv.h-1 ( when unit is loaded with its maximum capacity)
2.7.3 | Maximum leakage radiation level at 1 m from the source of the HDR unit in source 'OFF’ condition | ≤ 10 µSv.h-1 ( when unit is loaded with its maximum capacity)
2.8 | Treatment Planning system
2.8.1 | Digitizer and Localization Software
2.8.1.1 | Accuracy of digitization of Point co-ordinates | 3 % or 1mm
2.8.1.2 | Accuracy of data transfer from orthogonal radiograph | 3 % or 1mm
2.8.2 | Agreement between TPS and published/manual calculation of treatment time for single and multiple source loading/ dwell position, at relevant points | 3 % or 1mm
2.8.3 | Whether checklist of data ( e.g. source strength, unit of strength , distance and dose etc.) associated with TPS is in conjunction with quantity and units that will be used in clinical practice
2.8.4 | Agreement of source decay correction at different times of the day
2.8.5 | Agreement of source decay correction at different days of the week
2.8.6 | Correction for source strength decay before application
2.8.7 | Whether dose distribution is updated when dose calculation option and/or source characteristics are modified
2.8.8 | Quality of geometrical reconstruction (such as general shape, distance between marker and length of wires) using phantom of known geometry | 1 mm for short distance (< 2cm) and 2 mm for longer distances
2.9 | Safety Aspect in the Installation
2.9.1 | Whether Emergency Source Container provided
2.9.2 | Whether emergency procedures displayed near the control console
2.9.3 | Whether Radiation Warning symbol at the entrance provided
2.9.4 | Whether radiation levels all around the installtion found to be OK
2.9.5 | Whether Gamma Zone Monitor is approapriately installed & working