## Sheet1
DO NOT MODIFY THE WORK SHEET. FILL UP THE BLUE CELLS ONLY
Name of Medical Physicist responsible for the measurements
Date of completion of the measurements
1 | General information of the Telecobalt Units
1.1 | Manufacturer
1.2 | Model
1.3 | Sr. No.
1.4 | Maximum capacity of the source housing ( Ci/TBq)
1.5 | Current Source Strength as per source certificate ( Ci/TBq) as on date
1.6 | Source Loaded / replaced by (Name of the Certified Service Engineer)
1.7 | Total weight of Depleted Uranium in the unit, if any:
1.8 | Measured Source Strength in RMM (Exposure rate) /cGyMM (Air Kerma rate) as on date
1.9 | Radiation Output of the unit (cGy/min for 10x10 cm2 field at dmax, SSD Set up, measured using IAEA TRS 398 formalism)
1.1 | Available treatment modes (Fixed/Rotational/Arc/Skip)
1.11 | Field sizes available at isocenter (Min. and Max.)
1.12 | Normal Treatment Distance (NTD) (80 cm/ 100 cm)
1.13 | Secondary Collimator system (Symmetric / Asymmetric)
1.14 | Over travel distance for Asymmetric Jaws , if any
1.15 | Type of wedge filter ( Physical /Motorised/ Dynamic)
1.16 | Multileaf Collimator (MLC) ( available/not available)
1.16.1 | Type of MLC ( Secondary/Tertiary)
1.16.2 | If MLC is tertiary (Add on/ in-built)
1.16.3 | Minimum definable field size at isocentre
1.16.4 | Maximum definable field size at isocentre
1.16.5 | Number of leaves
1.16.6 | Leaf width at isocenter
1.17 | Any other technical information
2 | Performance test results of the Telecobalt unit
Sr. No. | Parameters | Tolerance Value, if any | Observation (Write Yes/No or give values of measurment as the case may be) | Observation is meeting the requirments | Remarks, if any
2.1 | Display and safety features
2.1.1 | Indicators for beam OFF (Green), beam ON & TRANSIT (Red) are provided on the source housing and working
2.1.2 | Head lock Indicator provided and working
2.1.3 | Radiation warning symbol provided on the outer surface of the Source Head
2.1.4 | Non electrical/ mechanical indicator on the Head for indicating Source ON/OFF and TRANSIT positions are provided and working
2.1.5 | Switches controlling the indicators (mentioned in sr. no. 2.1.4) are operated directly by the source drawer
2.1.6 | Source is automatically retrieved to OFF position in case of power failure
2.1.7 | Whether arrangement for manual retrieval of source in case it does not return to safe position is provided and working
2.1.8 | Limit on movement of major mechanical systems (Gantry, Collimator and Couch) in case of power failure:
2.1.8.1 | Rotational movements | ≤ 2o
21.8.2 | Linear movements | ≤ 10 mm
2.1.9 | Indicator lights on the Treatment Control Panel (TCP) provided and working: i) Ready state/Beam OFF – Green ii) source drawer in transit- RED iii) Radiation Beam ON- Yellow /Orange
2.1.10 | The switches to control the indicators (mentioned in sr. no. 2.1.9) are operated directly by the source drawer
2.1.11 | A key (or coded switch) on the TCP for proceeding to ready state is provided and working
2.1.12 | Irradiation can be started only by an operation from the TCP
2.1.13 | Irradiation and movements can be interrupted at any time from the TCP
2.1.14 | Whether provision to resume irradiation without pre-selection of operating parameters from TCP only are available and working (in case of interruption)
2.1.15 | Emergency switch on TCP to terminate irradiation is provided and working
2.1.16 | Emergency switches located at places other than TCP ( such as near gantry stand, patient support assembly and maze wall) are provided and working
2.1.17 | On termination, irradiation is not possible until all operating parameters are reselected from the TCP
2.1.18 | In the event of unplanned termination of irradiation (failure of timer/interlock etc) a display of the condition is provided and working at TCP
2.1.19 | Redundant/Primary-Secondary Timer combination to terminate irradiation are provided and working independently
2.1.20 | Switches to control the timer are actuated by the source drawer
2.1.21 | Pre-selected time is displayed until reset
2.1.22 | Elapsed time display is provided and working
2.1.23 | Whether mechanism to release the patient in case of power failure is provided and working
2.1.24 | In the event of power failure, information displayed at the time of failure is stored in one of the timers for at least 20 minutes
2.1.25 | Selection switches for stationary and moving beam radiotherapy are provided on the TCP and working
2.1.26 | Selection switch for CW/CCW rotation on TCP is provided and working
2.1.27 | Selection of specific wedge filter (or ZERO Filter) / tray is required at TCP to start irradiation
2.1.28 | Motorized Movements All motorized movements of the equipment (gantry, couch & collimators) possible only by personal actuation of two switches (enable switch) except during moving beam radiotherapy
2.2 | Couch
2.2.1 | Isocentre height above the floor level | ≤ 130 cm
2.2.2 | Isocentre clearance of the couch (with devices other than breast cone inserted) | ≥ 15 cm
2.2.3 | Couch isocentric rotation range | At least ±90 degree
2.2.4 | Lateral movement range of Table Top | ± 20 cm
2.2.5 | Vertical movements
2.2.5.1 | Minimum height from the floor level | ≤ 80 cm
2.2.5.2 | Vertical movement below the isocenter | ≥ 40 cm
2.2.5.3 | Vertical movement above the isocenter | ≥ 3 cm
2.2.6 | Longitudinal movement range of Table Top | ≥ 70 cm
2.2.7 | Minimum speed of linear motions of the couch | ≤ 1cm/sec
2.2.8 | Maximum speed of linear motions of the couch | ≤ 5cm/sec
2.2.9 | Minimum speed of rotational movement of Couch | ≤ 1o/sec
2.2.10 | Maximum speed of rotational movement of Couch | ≤ 7o/sec
2.2.11 | Couch Transmission (%) | Manufacturer Specified value :____ %
2.3 | Collimator
2.3.1 | Field size indicators are provided and working
2.3.2 | Collimator rotation range | At least ±900
2.4 | Gantry
2.4.1 | Gantry rotation range | At least ±1800
2.4.2 | Minimum speed of rotational movement of Gantry | ≤ 1o/sec
2.4.3 | Maximum speed of rotational movement of Gantry | ≤ 7o/sec
2.5 | Accessories
2.5.1 | Mechanism for securely latching of accessories to the equipment is provided and working
2.5.2 | Wedge filters are marked with their identity, angle and maximum field size
2.6 | Interlocks
2.6.1 | Interlock to prevent further irradiation provided in case of
2.6.1.1 | Failure of power supply to either of the timers
2.6.1.2 | Failure of primary timer to terminate irradiation
2.6.1.3 | One of the timers in a redundant timer combination is inoperative
2.6.1.4 | Source carrier does not attain BEAM ON/OFF condition within 3 sec of initiation /termination of irradiation
2.6.1.5 | Gantry moves during stationary beam radiotherapy
2.6.1.6 | Movement of the Gantry does not start/stop during Moving Beam radiotherapy
2.6.1.7 | Preselected angular movements are exceeded by more than 5o during moving beam radiotherapy
2.6.1.8 | Any selection in treatment room is different from the selection at TCP
2.6.1.9 | selected wedge/filter, if any, is not correctly positioned
2.6.1.10 | selected shielding tray, if any, is not correctly positioned
2.6.2 | Door interlock is provided and working
2.6.3 | Warning lights at the door to indicate Beam OFF, TRANSIT and ON positions is provided and working
2.6.4 | Audible alarm of irradiation in the treatment room is provided and working
2.7 | Mechanical and Radiation test
2.7.1 | Orthogonality of adjacent jaws | 900 ±1°
2.7.2 | Parallelism of opposite Jaws | ±1°
2.7.3 | Symmetry of opposite Jaws | ±1 mm
2.7.4 | Coincidence of optical beam and collimator axes | ±1 mm
2.7.5 | Accuracy of isocentre | 4 mm Dia sphere
2.7.6 | Accuracy of table top vertical motion | ±2 mm
2.7.7 | Accuracy of table top lateral motion | ±2 mm
2.7.8 | Accuracy of table top longitudinal motion | ±2 mm
2.7.9 | Accuracy of couch rotation scale | ±10
2.7.10 | Shift in optical field due to vertical motion of couch from min. to max. position | ±2 mm
2.7.11 | Shift in optical field due to rotational motion of the couch from -900 to 900 position | ±2 mm
2.7.12 | Accuracy of Collimator angular scale | ±10
2.7.13 | Accuracy of Optical Field Sizes for ≤ 10×10 cm2 | ±1 mm
2.7.14 | Accuracy of Optical Field Size for > 10×10 cm2 | ±2 mm
2.7.15 | Positional accuracy of asymmetric jaws | ±1 mm
2.7.16 | Optical Distance Indicator (ODI) range | SAD ± 20 cm
2.7.17 | Accuracy of optical distance indicator | ±2 mm
2.7.18 | Accuracy of Mechanical Front Pointer | ±2 mm
2.7.19 | Optical field overlap 00 /1800 | ±1 mm
2.7.20 | Optical field overlap 900 / 2700 | ±1 mm
2.7.21 | Optical & Radiation Field Congruence | ±2 mm
2.7.22 | Timer error | Manufacturer Specified value :____ sec
2.7.23 | Time taken by the source drawer to reach ON position and return back to OFF position | ≤ 5 sec
2.7.24 | Penumbra width for 10x10 cm2 field at dmax | ≤ 20 mm
2.7.25 | Coincidence between radiation and mechanical isocentre | ±2 mm
2.7.26 | Output constancy at different gantry angles | ±2%
2.7.27 | Beam uniformity (flatness) at gantry angles 00 and 900/2700 (For 10x10 cm2) at reference depth | ±3%
2.7.28 | Gantry Rotation speed (.1 to 1 RPM) | ±1%
2.7.29 | Accuracy of angular scale of gantry | ±0.50
2.7.30 | Tray position reproducibility | ±2 mm
2.7.31 | Table top sag at isocenter when loaded with 135kg distributed over 2m through isocentre | 2 mm
2.7.32 | Wedge factor for various wedges | Manufacturer Specified value_______
2.7.33 | Tray factor | Manufacturer Specified value:___
2.7.34 | Over travel distance for MLC, if available | Manufacturer Specified value_______
2.7.35 | Coincidence of MLC and secondary collimator axes | ±1 mm
2.7.36 | Reproducibility of leaf positions | ±1 mm
2.7.37 | Accuracy of Leaf positions | ±1 mm
2.7.38 | Congruence between MLC fields and beam’s eye view (BEV) | ±1 mm
2.7.39 | Coincidence between parallel opposed fields defined by MLC | ±1 mm
2.8 | Radiation Leakage/Transmission
2.8.1 | Transmission factor through beam Stopper, if any | 0.5%
2.8.2 | Relative Surface Dose at 0.5 mm depth (in percentage of absorbed dose at d max)
2.8.2.1 | For field size of 10X10 cm2 | ≤ 70%
2.8.2.2 | For largest field size available | ≤ 90%
2.8.3 | Maximum leakage radiation level at 5 cm from surface of the source housing in OFF position | 200 µSv/hr (when the unit is loaded with maximum capacity)
2.8.4 | Maximum leakage radiation level at 1 meter from the source in OFF position | 20 µSv/hr (when the unit is loaded with maximum capacity)
2.8.5 | *Maximum Transmission through beam limiting devices (Secondary collimators) | 2%
2.8.6 | *Maximum Leakage radiation through MLCs, if available | 2% (if secondary), 5% (if tertiary)
2.8.7 | *leakage radiation from the source housing in ON condition in the patient plane# | Maximum 0.2% Average 0.1%
2.8.8 | Maximum Leakage radiation from the source housing in ON condition at 1m from the source in other than patient plane | 0.5%
2.9 | Safety aspects of Telecobalt installation
2.9.1 | Whether emergency procedures displayed near the control console
2.9.2 | Laser provided and working
2.9.3 | Availability of T-Rod
2.9.4 | Radiation Warning symbol at the entrance provided
Note | *Percentage leakage radiation of maximum absorbed dose for 10 cm x 10 cm radiation field size measured on central beam axis at NTD when residual aperture is blocked with material of 2TVT
# Circular plane of radius 2 m centred on radiation beam axis at NTD